Table S1.  Summary statistics for PAH in sediments, Exercise 309 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH012MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Phenanthrene | 490 | 22 | 560 | 200 | 35 |
| Fluoranthene | 810 | 22 | 840 | 210 | 25 |
| Pyrene | 880 | 22 | 860 | 150 | 17 |
| Benz[a]anthracene | 480 | 22 | 490 | 120 | 25 |
| Chrysene | 480 | 22 | 490 | 150 | 30 |
| Benzo[b]fluoranthene | 690 | 21 | 780 | 240 | 31 |
| Benzo[e]pyrene | 680 | 16 | 650 | 120 | 18 |
| Benzo[a]pyrene | 660 | 21 | 650 | 160 | 25 |
| Benzo[ghi]perylene | 730 | 22 | 710 | 240 | 33 |
| Indeno[1,2,3-cd]pyrene | 650 | 20 | 660 | 200 | 30 |
|  |  |  | QPH013MS |  |  |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Phenanthrene | 930 | 22 | 950 | 260 | 28 |
| Fluoranthene | 1550 | 22 | 1480 | 270 | 18 |
| Pyrene | 1150 | 22 | 1090 | 200 | 19 |
| Benz[a]anthracene | 750 | 22 | 730 | 240 | 33 |
| Chrysene | 920 | 22 | 960 | 410 | 43 |
| Benzo[b]fluoranthene | 1140 | 21 | 1210 | 360 | 29 |
| Benzo[e]pyrene | 910 | 16 | 860 | 140 | 16 |
| Benzo[a]pyrene | 760 | 21 | 690 | 160 | 23 |
| Benzo[ghi]perylene | 750 | 22 | 690 | 180 | 26 |
| Indeno[1,2,3-cd]pyrene | 810 | 20 | 740 | 230 | 31 |

Table S2.  Summary statistics for PAH in sediments, Exercise 328 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH014MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Phenanthrene | 1390 | 22 | 1400 | 290 | 21 |
| Fluoranthene | 850 | 22 | 850 | 150 | 17 |
| Pyrene | 550 | 22 | 560 | 110 | 20 |
| Benz[a]anthracene | 260 | 22 | 270 | 70 | 25 |
| Chrysene | 310 | 22 | 310 | 90 | 28 |
| Benzo[b]fluoranthene | 270 | 22 | 270 | 80 | 28 |
| Benzo[e]pyrene | 180 | 17 | 210 | 90 | 43 |
| Benzo[a]pyrene | 180 | 22 | 180 | 40 | 22 |
| Benzo[ghi]perylene | 180 | 19 | 190 | 70 | 36 |
| Indeno[1,2,3-cd]pyrene | 180 | 20 | 190 | 50 | 28 |
|  |  |  | QPH015MS |  |  |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Phenanthrene | 1280 | 22 | 1310 | 290 | 22 |
| Fluoranthene | 2370 | 22 | 2380 | 550 | 23 |
| Pyrene | 2160 | 22 | 2260 | 580 | 26 |
| Benz[a]anthracene | 1180 | 22 | 1300 | 470 | 36 |
| Chrysene | 1480 | 22 | 1470 | 690 | 47 |
| Benzo[b]fluoranthene | 1410 | 22 | 1460 | 610 | 41 |
| Benzo[e]pyrene | 1340 | 17 | 1400 | 600 | 43 |
| Benzo[a]pyrene | 1170 | 22 | 1170 | 340 | 29 |
| Benzo[ghi]perylene | 1220 | 20 | 1340 | 680 | 51 |
| Indeno[1,2,3-cd]pyrene | 1120 | 20 | 1150 | 320 | 28 |

Table S3.  Summary statistics for PAH in sediments, Exercise 345 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH016MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 5 * | 12 | 5 | 2 | 48 |
| Acenaphthene | 13 | 19 | 13 | 6 | 48 |
| Fluorene | 28 | 20 | 29 | 13 | 44 |
| Naphthalene | 60 | 17 | 70 | 38 | 54 |
| Phenanthrene | 139 | 24 | 139 | 41 | 30 |
| Anthracene | 31 | 24 | 31 | 10 | 34 |
| Fluoranthene | 184 | 24 | 184 | 52 | 28 |
| Pyrene | 145 | 23 | 148 | 33 | 23 |
| Benz[a]anthracene | 77 | 23 | 79 | 24 | 31 |
| Chrysene | 101 | 23 | 103 | 25 | 24 |
| Benzo[b]fluoranthene | 133 | 20 | 141 | 46 | 32 |
| Benzo[k]fluoranthene | 71 | 21 | 75 | 37 | 49 |
| Benzo[e]pyrene | 102 | 18 | 105 | 27 | 26 |
| Benzo[a]pyrene | 87 | 23 | 85 | 29 | 35 |
| Dibenz[a,h]anthracene | 23 | 22 | 23 | 12 | 54 |
| Benzo[ghi]perylene | 97 | 24 | 97 | 28 | 29 |
| Indeno[1,2,3-cd]pyrene | 103 | 24 | 103 | 29 | 28 |

* entries in italics are indicative value only

Table S3 (continued).  Summary statistics for PAH in sediments, Exercise 345 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH017MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 24 | 12 | 24 | 15 | 65 |
| Acenaphthene | 113 | 19 | 113 | 50 | 44 |
| Fluorene | 347 | 20 | 347 | 137 | 39 |
| Naphthalene | 137 | 19 | 176 | 104 | 59 |
| Phenanthrene | 1440 | 23 | 1440 | 276 | 19 |
| Anthracene | 430 | 23 | 440 | 134 | 30 |
| Fluoranthene | 852 | 23 | 852 | 154 | 18 |
| Pyrene | 539 | 23 | 539 | 79 | 15 |
| Benz[a]anthracene | 247 | 23 | 247 | 42 | 17 |
| Chrysene | 304 | 23 | 304 | 58 | 19 |
| Benzo[b]fluoranthene | 272 | 23 | 273 | 87 | 32 |
| Benzo[k]fluoranthene | 141 | 21 | 146 | 50 | 34 |
| Benzo[e]pyrene | 202 | 18 | 202 | 38 | 19 |
| Benzo[a]pyrene | 172 | 23 | 168 | 48 | 29 |
| Dibenz[a,h]anthracene | 46 | 22 | 46 | 23 | 49 |
| Benzo[ghi]perylene | 165 | 23 | 165 | 23 | 14 |
| Indeno[1,2,3-cd]pyrene | 188 | 23 | 193 | 61 | 32 |

Table S4.  Summary statistics for PAH in sediments, Exercise 374 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH018MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 17.1 * | 13 | 39.7 | 37.7 | 95 |
| Acenaphthene | 35.6 | 17 | 37.9 | 23.7 | 62 |
| Fluorene | 65.2 | 17 | 65.3 | 32.0 | 49 |
| Naphthalene | 147.3 | 19 | 172.5 | 145.6 | 84 |
| Phenanthrene | 483.5 | 21 | 468 | 147 | 31 |
| Anthracene | 128.9 | 21 | 124.6 | 51.6 | 41 |
| Fluoranthene | 851.4 | 21 | 781 | 267 | 34 |
| Pyrene | 878.6 | 21 | 831 | 222 | 27 |
| Benz[a]anthracene | 448.6 | 21 | 411.0 | 180.3 | 44 |
| Chrysene | 448.8 | 21 | 433 | 176 | 41 |
| Benzo[b]fluoranthene | 612.8 | 21 | 629.2 | 284.1 | 45 |
| Benzo[k]fluoranthene | 367.8 | 20 | 406.4 | 203.8 | 50 |
| Benzo[e]pyrene | 596.1 | 17 | 584 | 101 | 17 |
| Benzo[a]pyrene | 656.5 | 21 | 605.4 | 223.0 | 37 |
| Dibenz[a,h]anthracene | 98.7 | 20 | 118.8 | 79.4 | 67 |
| Benzo[ghi]perylene | 595.2 | 21 | 557.6 | 255.5 | 46 |
| Indeno[1,2,3-cd]pyrene | 617.1 | 21 | 604 | 253 | 42 |

* entries in italics are indicative value only

Table S4 (continued).  Summary statistics for PAH in sediments, Exercise 374 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH019MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 49.3 * | 12 | 79.9 | 68.3 | 86 |
| Acenaphthene | 99.6 | 16 | 87.5 | 54.1 | 62 |
| Fluorene | 153.5 | 17 | 153 | 94.8 | 62 |
| Naphthalene | 872.0 | 18 | 688 | 506 | 73 |
| Phenanthrene | 947.5 | 21 | 894 | 308 | 34 |
| Anthracene | 278.6 | 21 | 261 | 95.1 | 36 |
| Fluoranthene | 1535 | 21 | 1463 | 454 | 31 |
| Pyrene | 1111 | 21 | 1045 | 301 | 29 |
| Benz[a]anthracene | 741.5 | 21 | 671 | 318 | 47 |
| Chrysene | 883.4 | 21 | 850 | 357 | 42 |
| Benzo[b]fluoranthene | 1035 | 21 | 1075 | 601 | 56 |
| Benzo[k]fluoranthene | 570.8 | 20 | 598 | 312 | 52 |
| Benzo[e]pyrene | 797.5 | 17 | 780 | 161 | 21 |
| Benzo[a]pyrene | 670.4 | 21 | 639 | 228 | 36 |
| Dibenz[a,h]anthracene | 131.4 | 19 | 166 | 117 | 70 |
| Benzo[ghi]perylene | 600.8 | 21 | 602 | 327 | 54 |
| Indeno[1,2,3-cd]pyrene | 648.5 | 21 | 648 | 313 | 48 |

* entries in italics are indicative value only

Table S5.  Summary statistics for PAH in sediments, Exercise 391 (assigned values, robust means and SDs in µg kg-1 dry weight).

|  |  |  | QPH020MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 6.9 | 8 | 6.9 | 5.4 | 79 |
| Acenaphthene | 12.7 | 10 | 14.5 | 10.9 | 75 |
| Fluorene | 21.3 | 13 | 24.6 | 20.2 | 82 |
| Naphthalene | 31.9 | 14 | 40.6 | 29.5 | 73 |
| Phenanthrene | 93.5 | 16 | 99.1 | 55.1 | 56 |
| Anthracene | 18.9 | 16 | 19.6 | 7.3 | 37 |
| Fluoranthene | 95.8 | 16 | 98.5 | 27.4 | 28 |
| Pyrene | 70.2 | 16 | 72.6 | 20.1 | 28 |
| Benz[a]anthracene | 39.6 | 16 | 41.4 | 15.6 | 38 |
| Chrysene | 44.6 | 16 | 46.1 | 14.9 | 32 |
| Benzo[b]fluoranthene | 53.2 | 15 | 58.0 | 17.7 | 30 |
| Benzo[k]fluoranthene | 32.6 | 15 | 39.4 | 22.5 | 57 |
| Benzo[e]pyrene | 42.3 | 15 | 44.1 | 13.4 | 30 |
| Benzo[a]pyrene | 36.7 | 16 | 37.8 | 11.7 | 31 |
| Dibenz[a,h]anthracene | 13.0 | 15 | 13.0 | 8.7 | 66 |
| Benzo[ghi]perylene | 34.2 | 16 | 35.4 | 12.6 | 36 |
| Indeno[1,2,3-cd]pyrene | 38.8 | 16 | 40.2 | 12.2 | 30 |

Table S5 (continued).  Summary statistics for PAH in sediments, Exercise 391 (assigned values, robust means and SDs in µg kg-1 wet weight).

|  |  |  | QPH021MS |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 5.6 | 7 | 5.6 | 5.8 | 104 |
| Acenaphthene | 9.3 | 9 | 11.0 | 9.3 | 84 |
| Fluorene | 14.8 | 13 | 17.3 | 17.5 | 101 |
| Naphthalene | 23.2 | 14 | 29.1 | 23.5 | 81 |
| Phenanthrene | 24.4 | 16 | 30.6 | 17.3 | 56 |
| Anthracene | 10.7 | 16 | 10.7 | 7.1 | 66 |
| Fluoranthene | 37.4 | 16 | 41.9 | 17.6 | 42 |
| Pyrene | 29.3 | 16 | 32.2 | 13.0 | 40 |
| Benz[a]anthracene | 17.0 | 16 | 20.0 | 8.9 | 45 |
| Chrysene | 21.8 | 16 | 22.7 | 9.1 | 40 |
| Benzo[b]fluoranthene | 27.5 | 15 | 29.7 | 9.6 | 32 |
| Benzo[k]fluoranthene | 17.5 | 15 | 18.7 | 9.1 | 49 |
| Benzo[e]pyrene | 19.2 | 15 | 20.4 | 6.6 | 32 |
| Benzo[a]pyrene | 18.1 | 16 | 19.3 | 7.6 | 39 |
| Dibenz[a,h]anthracene | 9.3 | 13 | 9.3 | 8.8 | 95 |
| Benzo[ghi]perylene | 16.6 | 16 | 17.3 | 7.7 | 44 |
| Indeno[1,2,3-cd]pyrene | 20.0 | 16 | 21.0 | 8.7 | 42 |

Table S6. Summary statistics for PAH in biological tissues, Exercise 394 (assigned values, robust means and SDs in µg kg-1  wet weight).

|  |  |  | QPH010BT |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 0.37 * | 8 | 2.42 | 3.68 | 152 |
| Acenaphthene | 0.89 | 11 | 3.10 | 3.12 | 101 |
| Fluorene | 1.73 | 11 | 3.45 | 3.23 | 94 |
| Naphthalene | 5.25 | 13 | 8.92 | 6.85 | 77 |
| Phenanthrene | 12.0 | 16 | 15.2 | 6.93 | 46 |
| Anthracene | 1.58 | 14 | 1.87 | 1.22 | 65 |
| Fluoranthene | 31.4 | 16 | 33.0 | 7.37 | 22 |
| Pyrene | 23.1 | 16 | 24.2 | 5.84 | 24 |
| Benz[a]anthracene | 10.6 | 16 | 12.2 | 4.54 | 37 |
| Chrysene | 17.3 | 16 | 19.0 | 8.20 | 43 |
| Benzo[b]fluoranthene | 13.3 | 14 | 14.6 | 5.02 | 34 |
| Benzo[k]fluoranthene | 5.08 | 13 | 9.68 | 8.80 | 91 |
| Benzo[e]pyrene | 13.2 | 15 | 14.1 | 3.88 | 28 |
| Benzo[a]pyrene | 5.57 | 15 | 6.38 | 2.83 | 44 |
| Dibenz[a,h]anthracene | 1.23 | 13 | 2.55 | 2.67 | 105 |
| Benzo[ghi]perylene | 5.47 | 14 | 5.97 | 1.89 | 32 |
| Indeno[1,2,3-cd]pyrene | 4.46 | 14 | 5.83 | 3.19 | 55 |

Table S6 (continued). Summary statistics for PAH in biological tissues, Exercise 394 (assigned values, robust means and SDs in µg kg-1  wet weight).

|  |  |  | QPH011BT |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Assigned value (µg kg-1) | Number of observations | Robust mean (µg kg-1) | Robust SD between laboratories | Robust CV% |
| Acenaphthylene | 1.60 | 9 | 3.30 | 3.20 | 97 |
| Acenaphthene | 95.9 | 11 | 112 | 49.0 | 44 |
| Fluorene | 60.1 | 13 | 64.8 | 15.7 | 24 |
| Naphthalene | 92.0 | 13 | 109 | 64.1 | 59 |
| Phenanthrene | 68.4 | 16 | 72.1 | 15.5 | 21 |
| Anthracene | 7.56 | 16 | 9.53 | 4.83 | 51 |
| Fluoranthene | 33.8 | 16 | 35.5 | 6.68 | 19 |
| Pyrene | 19.8 | 16 | 20.4 | 7.78 | 38 |
| Benz[a]anthracene | 5.08 | 16 | 6.07 | 2.65 | 44 |
| Chrysene | 6.69 | 16 | 7.14 | 2.24 | 31 |
| Benzo[b]fluoranthene | 5.88 | 14 | 6.98 | 3.74 | 54 |
| Benzo[k]fluoranthene | 2.54 | 13 | 4.15 | 3.50 | 84 |
| Benzo[e]pyrene | 5.52 | 15 | 5.94 | 1.75 | 29 |
| Benzo[a]pyrene | 3.09 | 14 | 3.76 | 2.61 | 69 |
| Dibenz[a,h]anthracene | 3.96 | 8 | 3.96 | 6.79 | 171 |
| Benzo[ghi]perylene | 2.03 | 12 | 2.25 | 0.56 | 25 |
| Indeno[1,2,3-cd]pyrene | 2.31 | 13 | 3.32 | 2.41 | 73 |

* entries in italics are indicative value only

Table S7.  Summary of Laboratory Performance: numbers of analyses made by each laboratory have been split by analytical method. Laboratories have been ranked in descending order by the percentage of satisfactory results.

| Laboratory Code | No. of Obs. | No. of Obs. by Method |  | % Satisfactory | % Questionable | % Unsatisfactory | % Extremely Unsatisfactory |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | GC | HPLC |  |  |  |  |
| Q311A | 29 |  | 29 | 100 |  |  |  |
| Q138C | 109 |  | 109 | 97 | 3 |  |  |
| Q107A | 40 |  | 40 | 95 | 5 |  |  |
| Q357A | 19 | 19 |  | 95 | 5 |  |  |
| Q115A | 91 | 71 | 20 | 93 | 5 | 1 |  |
| Q143A | 78 |  | 78 | 91 | 4 | 5 |  |
| Q251A | 31 | 31 |  | 90 | 6 | 3 |  |
| Q131D | 40 | 40 |  | 90 | 5 | 5 |  |
| Q114A | 140 |  | 140 | 88 | 7 | 4 | 1 |
| Q152A | 195 | 195 |  | 88 | 4 | 8 |  |
| Q127A | 45 |  | 45 | 87 | 7 | 4 | 2 |
| Q111B | 65 |  | 65 | 86 | 9 | 5 |  |
| Q134B | 20 |  | 20 | 85 | 5 | 5 | 5 |
| Q118A | 91 |  | 91 | 80 | 19 | 1 |  |
| Q126C | 149 |  | 149 | 80 | 14 | 6 |  |
| Q210A | 117 | 117 |  | 79 | 12 | 7 | 2 |
| Q104C | 46 | 46 |  | 74 | 7 | 13 | 7 |
| Q131B | 161 | 161 |  | 71 | 11 | 16 | 2 |
| Q275A | 79 | 79 |  | 67 | 9 | 16 | 8 |
| Q108A | 60 | 48 | 12 | 65 | 17 | 18 |  |
| Q116A | 46 | 46 |  | 63 | 15 | 22 |  |
| Q130B | 70 | 70 |  | 63 | 19 | 16 | 3 |
| Q121A | 46 |  | 46 | 61 | 20 | 11 | 9 |
| Q153A | 182 | 164 | 18 | 60 | 21 | 14 | 4 |
| Q114D | 94 |  | 94 | 57 | 15 | 18 | 10 |
| Q110A | 40 | 40 |  | 55 | 13 | 15 | 18 |
| Q205A | 44 | 44 |  | 55 | 16 | 23 | 7 |
| Q234A | 126 | 126 |  | 54 | 2 | 3 | 41 |
| Q122A | 102 | 102 |  | 51 | 22 | 10 | 18 |
| Q224A | 89 | 89 |  | 49 | 16 | 26 | 9 |
| Q157A | 37 | 37 |  | 46 | 22 | 30 | 3 |
| Q113B | 82 |  | 82 | 45 | 9 | 7 | 39 |
| Q134D | 51 |  | 51 | 45 | 27 | 22 | 6 |
| Q152C | 18 | 18 |  | 44 | 28 | 28 |  |
| Q240A | 78 | 78 |  | 42 | 5 | 29 | 23 |
| Q128A | 59 | 59 |  | 41 | 10 | 12 | 37 |
| Q134A | 125 | 71 | 54 | 38 | 10 | 38 | 14 |
| Q250A | 10 |  | 10 | 30 |  | 50 | 20 |
| Q204A | 20 | 20 |  | 25 | 10 | 40 | 25 |
| Q226A | 79 | 79 |  | 8 | 16 | 24 | 52 |
| Q242A | 1 | 1 |  | 0 | 100 |  |  |